XRZ-10000

T.E.N.S. Unit

Cooke Inc.

User’s Manual

Thank you for choosing the XRZ-10000 T.E.N.S. unit.

Please read this instruction manual for safety and proper use of this unit.

Please keep this manual for future reference.

Table of contents:

- Safety

          - Cautions

          - Warnings

Features

- Specifications

- Components

- How to use the unit

- Schematic/Board

- Programming

- Troubleshooting

- Technician Troubleshooting

- Appendix

Safety

Cautions:

- Read instruction manual before using this unit.

- Please consult a physician or other qualified personnel before using TENS.

- Isolated cases of skin irritation may occur at the site of electrode placement following long term application.

- Possible adverse reaction to gel, or electrodes may occur in some users. If you have a known allergy then please consult a physician before use.

- We emphasize that users with an implanted electronic device (for example, a pacemaker) should not use TENS treatment without first consulting a physician.

- If TENS therapy becomes ineffective, unpleasant, or you experience any adverse effects then use should be discontinued immediately until you consult with a physician or other qualified personnel.

WARNING:

- Do not use if you have an electrical implanted device without first consulting your physician or other qualified personnel.

- Do not use if you have been diagnosed with any heart conditions (for example: heart murmur, High Blood Pressure, Cardiomyopathy) before consulting a physician or other qualified personnel.

- Do not use if you during pregnancy or birth without first consulting a physician or other qualified personnel.

- Do not use on your neck or head (for example, do not place electrodes on or over eyes, mouth, throat or carotid artery) as doing so can result in injury or death.

- Do not operate or drive machinery or vehicles while using TENS.

- Do not use in or around explosive or flammable materials or vapors.

- Keep this device out of reach from children.

Features

- Easy to use unit and completely user friendly

- Easy to read display

- Offers the user complete control of the frequency and intensity with easy to use dial controls

- Three different modes that can be change by the push of a button

- Wall plug in so there is no need for buying or replacing batteries

Specifications

Max input: 120V AC maximum input voltage

Max outputs: Electrodes: 117V DC maximum @ 64μ

Max Power consumption: 7W maximum

8

1

7

6

5

4

3

Construction

2

- Power Switch

- Channel 1

- Channel 2

- Frequency adjustment

- Amplitude adjustment

- LCD

- Waveform Selector Button

- Reset Button

How to use this unit

- Please consult a physician or other qualified personal before using this TENS unit.

- First make sure that electrodes are not placed in contact with skin before plugging the unit in.

- Then make sure the power is not switch on. Once verified that the unit power switch is not in the on position you can plug the wall wart in to the wall outlet.

- Now check to make sure that all adjustment knobs are turned to their minimum setting.

- Plug electrode jacks into channels 1 & 2

- Next turn on the unit.

- You may now apply the electrodes to your desired application area specified by your physician or other qualified personal.

- Once on you may turn the adjustment knobs and change the waveform to your desired setting by pressing the waveform selector button in order to change waveforms (first press is square wave, second press is sinewave, and third press is triangle wave.

- You can hit the reset button in order to immediately go to the square wave function at any time.

Troubleshooting

- If no stimulation is felt from the electrodes these steps should be taken. Make sure to set the adjustment to the lowest setting before troubleshooting.

- First make sure you have the unit plug into the wall and that the power button is in the ON position.

- Next you should make sure that you have a solid and clean contact from electrode to skin. If not replace the electrodes with new ones. If you replaced the electrodes then increase the amplitude until stimulation is felt. If no stimulation is present continue to step 3.

- Remove the electrodes from your skin and turn the power off to the unit then unplug the unit from the wall.

- Put in a cardboard box and ship to 124 Admin Drive, Randolph Center, VT 05061. Where we will fix it or send you a new one.*

*For $59.99 plus shipping and handling. Tax not included.

Technician Troubleshooting

- There should be a square waveform between 0-6Vpeak at the test point shown in figure 1.

- There should be either a sine or triangle waveform between 0-6Vpeak at the test point shown in figure 1.1.

- There should be 5V at the output of the LM7805 as shown in figure 1.2.

- There should be a waveform between 35V-188V peak to peak at the output of the transformer as shown in figure 1.3.

- At pins 7 and 8 there should be 3V with respect to ground.

- There should be 6V at the voltage divider with respect to ground.

- Relay 1 is activated by applying a positive voltage to the base of the transistor which is connected to pin E23 on the Freescale board.

- Relay 2 is activated in the same way as Relay 1 but the base of its transistor is connected to pin E22 on the Freescale board.

- When the WaveformState is equal to 1 in the code then relay 1 should be deactivated and relay 2 should be activated. This should produce a square wave at the input of the audio amplifier and produce a square wave at the output of the transformer.

- When the WaveformState is equal to 2 in the code then Relay 1 should be deactivated and relay 2 should also be deactivated. This should produce a square wave at the input of the audio amplifier and at the output of the transformer.

- When the WaveformState is equal to 2 in the code then Relay 1 should be activated and relay 2 should be deactivated. This should produce a triangle wave at the input of the audio amplifier and at the output of the transformer.

- Frequency is equal to fo=1RC  R equals the resistance at pin 7 and C equals the capacitor value at pin 5 to 6. With R equaling 3K and C equaling .47μF the frequency is equal to 709Hz.

- Frequency is adjust with the POT at pin 7.

- Amplitude is proportional to the resistance at pin 3 when connected to the middle of the voltage divider. For every KΩ at connected pin 3 an increase of 60mV at the output for sine and square, 160mV for triangle.

- Amplitude is adjusted with the POT at pin 3.

System Diagram

Sine/Triangle wave test point

Figure 1.1

Square wave test point

Figure 1

Transformer Test Points

5V Test Point

Figure 1.2

Figure 1.3

Schematic:

Board Layout:

Programming

/*

* LCDTENS2.c  Created on: Apr 24, 20 :Author: kdc07170

*/

#include "derivative.h" /* include peripheral declarations */

#include "my_io_apps.c" /* include peripheral declarations */

#include "mcg.h" /* using XTAL and PLL does not work returns errors*/

#include "my_interrupts.c"

#include "my_lcd2.c" /* include lcd declarations */

/**************Global Variables Here************************/

int WaveformState = 0; //1,2,3 for square, sine, triangle

#define RelaysquareON gpo_e(22,1) //setting gpo_e(22,1) as relaysquareON

#define RelaysquareOFF gpo_e(22,0) //setting gpo_e(22,0) as relaysquareOFF

#define RelaySineON gpo_e(23,1) //setting gpo_e(23,1) as relaysineON

#define RelaySineOFF gpo_e(23,0) //setting gpo_e(23,0) as relaysineOFF

#define SelectPB gpi_c(2) //setting gpi_c(2) as SelectPB

#define ResetPB gpi_b(3) //setting gpi_b(3) as ResetPB

#define Freqcounter gpi_c(1) //setting gpi_c(1) as Freqcounter

int TimeStamp1, TimeStamp2;//declaring timestamp 1 & 2 as a integers

float Frequency;//declaring frequency as a floating integer

void displayfreq(void) ;//creating display freq as a subroutine

int main(void)

{

//*******Local Variable Declarations for main *****************

//******Initial Chip Settings******************

int pll_err=pll_init(8000000,0,1,4,24,1);     //Use the crystal oscillator to drive the PLL

//******User chip initializations***************************************

init_gpio_e(22,1);//initialization relays

init_gpio_e(23,1);

init_gpio_d(2,0);// initialization Frequency counter

init_gpio_c(2,0);// initialization pushbuttons(2)

init_gpio_b(3,0);

//Initialize the lcd

LCD_init(DISPLAY_8X5|_2_LINES,DISPLAY_ON|CURSOR_OFF|CURSOR_BLINK);

syst_tick_init();

// Convert the specific pins to the GPIO module.  Pins B18,B19, D1

LCD_write_string("Freq:"); //writing to the LCD

LCD_pos_xy(0,1); //changing postion of cursor on LCD

LCD_write_string("Waveform:"); //writing to the LCD

LCD_pos_xy(20,0); //changing postion of cursor on LCD

for(;;)

{

LCD_pos_xy(1,0); //changing postion of cursor on LCD

sys_ticks=0; //setting internal timer sys_ticks to 0

if(WaveformState==4) //checking to see if WaveformState is 								equal to 4

{WaveformState=0;} //If WaveformState is equal to 4 then 									WaveformState is reset to 0

if(SelectPB==0 && WaveformState<4) //checking to see if                    SelectPB is pressed and that WaveformState is less than 4

{

while(SelectPB ==0); WaveformState=WaveformState+1;     				//If SelectPB is pressed then do wait for it to be released and then increament WaveformState by 1

if(WaveformState==1) //If waveformstate equals 1 then turn          		          write to LCD and output squarewave

{LCD_write_string("Square ");RelaysquareON;}

if(WaveformState==2) //If waveformstate equals 2 then turn          		                  write to LCD and output sinewave

{RelaySineON;RelaysquareOFF;LCD_write_string("Sine    ");}

if(WaveformState==3) //If waveformstate equals 3 then turn          		         write to LCD and output trianglewave

{RelaySineOFF;RelaysquareOFF;LCD_write_string("Triangle");}

}

if(ResetPB==0){WaveformState=1;} //If resetPB is pressed then reset Waveformstate to 1

//displayfreq();//under construction

}

return 0;

}

void displayfreq(void)

{

sys_ticks = 0;

while (ECHO==0);

TimeStamp1= sys_ticks;

while (ECHO==1);

TimeStamp2= sys_ticks;

Frequency= 1/(2*(TimeStamp1-TimeStamp2));

LCD_pos_xy(10,0);

sprintf(buffer, "%3i", Frequency);

}

Appendix